## High volume BCR error
Smart Center - EC Proactive analysis |  | General Algorithm
Instrument : A3600
Subsystem : IOM - Barcode Reader v. 1.0 - 11/20/2015
EC(s): UPS: 0405 - Unreadable Sample ID or Unreadable Rack ID
Orginator: Rusty F. Clay
Failure mode
0405 - Sample Tube cannot be processed and is sent to Priority Output Rack because Barcode ID is not readable
*Low volume days are eliminated from equaition to isolate outreach samples vs in-house samples. The purpose of this is to determine if bad labels are the cause of the errors and to assist in isolating where the source (in-house or outreach).
Limitations of process :
Connectivity to AbbottLink
False flag conditions : A- During preventative maintenance or a repair if the FSR/CLS is performing a PM on that particular module the number of errors for that day may be falsly elevated. A two consecutive day measure is used to eliminate a false flag condition.
Failure mode examples and suggested evaluations:
Data file(s) format
layout.csv file
| canID | canAddr | nodeType | description | initialize | nextNodeID | jumpNodeID | nodePath | instance | equivalentNodes
| 001 | 001 | IOM | Input/Output | 1 | 002 |  |  | 1
| 002 | 002 | CM | Centrifuge | 1 | 003 |  |  | 1 | 004 005
| 003 | 003 | TTM | Track T Module | 1 | 004 | 005 | 4 5 | 1
| 004 | 004 | CM | Centrifuge | 1 | 005 |  |  | 2 | 005
| 005 | 005 | CM | Centrifuge | 1 | 006 |  |  | 3
| 006 | 006 | DCM | Decapper | 1 | 007 |  |  | 1 | 007
| 007 | 007 | DCM | Decapper | 1 | 008 |  |  | 2
|  |  |  | description of module | [5 empty] | indicates more than one module
| canID is the equivalent of nodeID in error log | [6 empty] | module instance
error log file
Log file type
15 = Year, 11 = Month, 30 = Day, 23 = hour, 59 = minute
ERROR - 1511302359.XML
<Error timestamp="20151210050523" nodeID="1" nodetype="IOM" errorcode="0405" sampleID="&amp;U42" moreinfo="Unreadable Sample ID or Unreadable Rack ID " Off-line="No"/>
<Error timestamp="20151210050523" nodeID="1" nodetype="IOM" errorcode="0405" sampleID="" moreinfo="Unreadable Sample ID or Unreadable Rack ID " Off-line="No"/>
<Error timestamp="20151210050526" nodeID="1" nodetype="IOM" errorcode="0405" sampleID="&amp;U" moreinfo="Unreadable Sample ID or Unreadable Rack ID " Off-line="No"/>
|  |  | Time of error | NodeID |  |  | Error Code |  |  | Additional error message text
| Node = Module |  | Tubes Today = Number of tubes processed (tubes/day) by module
Counter-20151201.txt
| Node | Gates Today | Gates Prev Days | Tubes Today | Tubes Prev Days | Unloaded Today | Unloaded Prev Days | Disposed Today | Disposed Prev Days
| TRU-1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| IOM-1 | 79593 | 46225619 | 1271 | 1273752 | 176 | 346720 | 0 | 0
| CM-1 | 79593 | 46255293 | 145 | 151354 | 146 | 151698 | 0 | 0
| TTM-1 | 79597 | 46259938 | 0 | 0 | 0 | 0 | 0 | 0
| CM-2 | 13252 | 10310998 | 145 | 164626 | 147 | 165396 | 0 | 0
| CM-3 | 13250 | 10310974 | 145 | 162497 | 143 | 162979 | 0 | 0
Module Descriptor
A3600 files collected from IDA 01:00 local time (error and layout)
Barcode Error % Calculation: Count the number of 0405 errors per IOM Module.
divide that number by the number of tubes at IOM-"X". Compare that number to the previous 30 days % of error. Additionally ignore any days that have less than 1 SD of 30 day average tubes/day. This gives us "high volume % Error"
Do any of the error codes exceed 1.5 SD % error for 2 consecutive days?
Done
Yes
No
Generate Alert and send to CLS/FSR
NodeID identifies different/separate modules.

## Low Volume BCR error
Smart Center - EC Proactive analysis |  | General Algorithm
Instrument : A3600
Subsystem : IOM - Barcode Reader v. 1.0 - 11/20/2015
EC(s): UPS: 0405 - Unreadable Sample ID or Unreadable Rack ID
Orginator: Rusty F. Clay
Failure mode
0405 - Sample Tube cannot be processed and is sent to Priority Output Rack because Barcode ID is not readable
*High vilue days are eliminated from equaition to isolate in-house samples vs outreach samples. The purpose of this is to determine if bad labels are the cause of the errors and to assist in isolating where the source (in-house or outreach).
Limitations of process :
Connectivity to AbbottLink
False flag conditions : A- During preventative maintenance or a repair if the FSR/CLS is performing a PM on that particular module the number of errors for that day may be falsly elevated. A two consecutive day measure is used to eliminate a false flag condition.
Failure mode examples and suggested evaluations: | [5 empty] | Note: 2 consecutive low volume days are consecutive low volume days. This may be Sat Sun Sat Sun, because Mon - Fri were high volume days. Example Sun to next Sat would be low volume consecutive days.
Data file(s) format
layout.csv file
| canID | canAddr | nodeType | description | initialize | nextNodeID | jumpNodeID | nodePath | instance | equivalentNodes
| 001 | 001 | IOM | Input/Output | 1 | 002 |  |  | 1
| 002 | 002 | CM | Centrifuge | 1 | 003 |  |  | 1 | 004 005
| 003 | 003 | TTM | Track T Module | 1 | 004 | 005 | 4 5 | 1
| 004 | 004 | CM | Centrifuge | 1 | 005 |  |  | 2 | 005
| 005 | 005 | CM | Centrifuge | 1 | 006 |  |  | 3
| 006 | 006 | DCM | Decapper | 1 | 007 |  |  | 1 | 007
| 007 | 007 | DCM | Decapper | 1 | 008 |  |  | 2
|  |  |  | description of module | [5 empty] | indicates more than one module
| canID is the equivalent of nodeID in error log | [6 empty] | module instance
error log file
Log file type
15 = Year, 11 = Month, 30 = Day, 23 = hour, 59 = minute
ERROR - 1511302359.XML
<Error timestamp="20151210050523" nodeID="1" nodetype="IOM" errorcode="0405" sampleID="&amp;U42" moreinfo="Unreadable Sample ID or Unreadable Rack ID " Off-line="No"/>
<Error timestamp="20151210050523" nodeID="1" nodetype="IOM" errorcode="0405" sampleID="" moreinfo="Unreadable Sample ID or Unreadable Rack ID " Off-line="No"/>
<Error timestamp="20151210050526" nodeID="1" nodetype="IOM" errorcode="0405" sampleID="&amp;U" moreinfo="Unreadable Sample ID or Unreadable Rack ID " Off-line="No"/>
|  |  | Time of error | NodeID |  |  | Error Code |  |  | Additional error message text
| Node = Module |  | Tubes Today = Number of tubes processed (tubes/day) by module
Additional considerations need to be taken. One barcode read error actually generates more than one error message. The time at which this error occurs is the determining factor of how many errors occur. In this instance since all of the 0405 errors occur within a timespan of 3 seconds, this indicates a single barcode read error, not 3.
Log file type
2015 = Year, 12 = Month, 01 = Day
Counter-20151201.txt
| Node | Gates Today | Gates Prev Days | Tubes Today | Tubes Prev Days | Unloaded Today | Unloaded Prev Days | Disposed Today | Disposed Prev Days
| TRU-1 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
| IOM-1 | 79593 | 46225619 | 1271 | 1273752 | 176 | 346720 | 0 | 0
| CM-1 | 79593 | 46255293 | 145 | 151354 | 146 | 151698 | 0 | 0
| TTM-1 | 79597 | 46259938 | 0 | 0 | 0 | 0 | 0 | 0
| CM-2 | 13252 | 10310998 | 145 | 164626 | 147 | 165396 | 0 | 0
| CM-3 | 13250 | 10310974 | 145 | 162497 | 143 | 162979 | 0 | 0
Module Descriptor
A3600 files collected from IDA (error and layout)
Barcode Error % Calculation: Count the number of 0405 errors per IOM Module.
divide that number by the number of tubes at IOM-"X". Compare that number to the previous 30 days % of error. Additionally ignore any days that have greater than 1 SD of 30 day average tubes/day. This gives us "low volume % Error" NOTE
Do any of the error codes exceed 1.5 SD % error for 2 consecutive days?
Done
Yes
No
Generate Alert and send to CLS/FSR
NodeID identifies different/separate modules.

## Sheet2
IHN |  | Type of Error | Module | Error Code | Error Description | Threshold
2 | A36-01 | Barcode Reader Errors | IOM | 0405 | Unreadable Sample ID or Unreadable Rack ID | 1 SD above 30 day mean for 2 consecutive days (high volume)
2 | A36-02 | Barcode Reader Errors | IOM | 0405 | Unreadable Sample ID or Unreadable Rack ID | 1SD above 30 day mean for 2 consecutive days (low volume)
2 | A36-03 | Barcode Reader Errors | CM | 0605 | Unreadable Sample ID | 1.5%/day for 2 consecutive days
2 | A36-03 | Barcode Reader Errors | SRM | 0E05 | Unreadable Sample ID | 1.5%/day for 2 consecutive days
2 | A36-03 | Barcode Reader Errors | ISR | 5015 | Unreadable Sample ID | 1.5%/day for 2 consecutive days
2 | A36-03 | Barcode Reader Errors | AQM | E044 | Unreadable Sample ID | 1.5%/day for 2 consecutive days
2 | A36-04 | Barcode Reader Errors | ALL | SC05F | Too Many Sample ID Mismatches | 1/day for 2 consecutive days